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REVISION HISTORY
	Version
	Author
	Summary of change

	00
	Emma Pierce
	First document for approval.

	01
	James J. Stors
	New software version to print serial number in the Report footer.

	[bookmark: _Toc442432502]02
	Luo HuaZhen
	New software version to add Optics Bench part numbers


Purpose
The 2998 Final Test System, TE82882-8, is located in Singapore at contract manufacturer Flextronics.  The 2998 Final Test System is used to perform functionality testing of Waters Corporation’s 2998 Photo Diode Arrays (PDA’s) as defined in TSP 735001613, to ensure compliance to the defined Manufacturing Test Specifications.
The system uses serial ports to communicate to up to four (4) UUT’s in order to test Noise, Drift, and Optics. Upon completion a Final Test Record is automatically generated and a hard copy is printed. In addition, the Final Test Record is sent to the SigmaProd Server.
[bookmark: _Toc442432503]Scope of Report
This Test Equipment Report details the IQ activities undertaken for the 2998 Final Test System (TE82882-8) based on the requirements of the performed Installation Qualification, 730004135, Version 02 as outlined by the Test Software Version 4.2 risk assessment, per TE82882-10JUN19-VAF.
The risk assessment, based on this change, resulted in the requirement of performing the following test scripts:
· IQ-01 Associated Documentation Verification: To verify that the PDA Final Test System Test and Specification Procedure, 735001613, that references the new software version, 4.2, is in the appropriate state in SAP.
· IQ-04 Test Software Version: To verify the version specified within the PDA Final Test, Test and Specification Procedure, 735001613, matches the version of the program.
· IQ-06 Functional Verification: To verify that the entered UUT Serial Number is correctly listed at the bottom of the Test Record.
All other testing performed, as detailed in TE82882-8-IQR, Iteration 00, remains valid and applicable as there is no change to the installation site, calibration and PM schedule, or the safety and alarms.
[bookmark: _Toc442432509]Summary of Results - Statement of Fitness for Intended Use 
The 2998 Final Test System (TE82882-8) was tested in accordance with the applicable sections of the Installation Qualification Protocol, 730004135, Version 02, as outlined in the Test Software Version 4.2 risk assessment, per TE82882-10JUN19-VAF. All required Test Scripts were performed and all objective evidence requirements have been met.
Waters now has the necessary documentation in place to ensure that the 2998 Final Test System (TE82882-8) remains compliant with Waters’ requirements, the latest relevant quality system regulations, and cGMP practices. 
Following the approval of this Qualification Report, the 2998 Final Test System (TE82882-8) remains fully qualified and fit for intended use post the change, as outlined in TE82882-10JUN19-VAF.
Glossary
The following definitions are used in this document:
	Term
	Explanation

	cGMP
	Current Good Manufacturing Practice

	IQ
	Installation Qualification

	IQR
	Installation Qualification Report

	UUT
	Unit Under Test

	PDA
	Photo Diode Array

	SAP
	Enterprise Resource Planning (ERP) software

	TSP
	Test Specification and Procedure

	PM
	Preventive Maintenance

	VAC
	Volts Alternating Current

	Hz
	Hertz

	ºC
	Degrees Celsius

	N/A
	Not Applicable


[bookmark: _Toc442432507]Training
[bookmark: _Toc442432506]The Executor who performed the 2998 Final Test System (TE82882-8) Installation Qualification was trained according to the Installation Qualification Procedure, 730004135, Version 02. The training record is proved as Attachment 7.1.
The Operator who performed the testing of the UUT using the 2998 Final Test System (TE82882-8) per Installation Qualification Procedure, 730004135, Version 02, IQ-06, Functional Verification Test Script, has been trained to the updated PDA Final Test System Test and Specification Procedure, 735001613, Version 22.  The Training Record is provided as Attachment 7.2.
Qualification Results
[bookmark: _Toc481749069]Test Script IQ-01, Associated Documentation Verification, was performed in accordance with Installation Qualification Protocol, 730004135, Version 02, for the PDA Final Test System Test and Specification Procedure, 735001613, as outlined in TE82882-10JUN19-VAF.
[bookmark: _Toc481749070]IQ-01, Associated Documentation Verification Data Sheet, is provided as Attachment 7.3.
[bookmark: _Toc481749071]The updated PDA Final Test System Test and Specification Procedure, 735001613, Version 22, was confirmed to have been updated with software version 4.2 and to be in SAP with a status of Ready. A screen capture is provided below: 
[image: ]
The PDA Linearity Test System Preventive Maintenance Procedure, 735003030, Version 00/Revision A, was confirmed to be in SAP with a status of Release General. A screen capture is provided below:
[image: ]



[bookmark: _Toc481749091]Test Script IQ-04, Test Software and Instrument Firmware Version, was performed in accordance with Installation Qualification Protocol, 730004135, Version 02, for the updated software version 4.2, as outlined in TE82882-10JUN19-VAF.
[bookmark: _Toc481749092]IQ-04, Test Software and Instrument Firmware Version Data Sheet, is provided as Attachment 7.4.
The Test Equipment Software version listed in 735001613, PDA Final Test System Test and Specification Procedure (Version 22), is 4.2 and was verified to match the revision (listed as Revision on the Operator Interface) displayed on the Test Equipment PDA Final Test System TE82882 program display, as shown in the screen capture below:
[image: ]
[bookmark: _Toc481749098]Test Script IQ-06, Functional Verification, was performed in accordance with Installation Qualification Protocol, 730004135, Version 02, for the updated software version 4.2, as outlined in TE82882-10JUN19-VAF.
[bookmark: _Toc481749099]IQ-06, Functional Verification Data Sheet is provided as Attachment 7.5.
One (1) UUT, Part Number 700012243, was tested using the TE82882-8, 2998 Final Test System, and met the requirement of obtaining a Final Test Record with the UUT Serial Number correctly listed at the bottom of the Test Record.
The Instrument Final Test Record is provided as Attachment 7.6.
One (1) UUT, Part Number 700012245, was tested using the TE82882-8, 2998 Final Test System, and met the requirement of obtaining a Final Test Record with the UUT Serial Number correctly listed at the bottom of the Test Record.
The Instrument Final Test Record is provided as Attachment 7.7.

A summary of the attached objective evidence that applies to each Test Script is provided in the Table below:
	Script or Section
	Pass/Fail
	Objective Evidence
	Deviations

	IQ-01
	Pass
	Attachment 7.3
	None

	IQ-04
	Pass
	Attachment 7.4
	None

	IQ-06
	Pass
	Attachment 7.5, 7.6 and 7.7
	None


[bookmark: _Toc442432505][bookmark: _Toc309978937][bookmark: _Toc442432504]Attachments
0.   Completed Executor Training Record 730004135, Ver.02
0.   Completed Operator Training Record 735001613, Ver.22
0.2 IQ-01, Associated Documentation Verification Test Script Data Sheet 
0.3 IQ-04, Test Software and Instrument Firmware Version Test Script Data Sheet 
0.4 IQ-06, Functional Verification Test Script Data Sheet
0.5 Instrument Final Test Record for IQ-06, Part Number: 700012243, Serial Number: 22K6P003
0.6 Instrument Final Test Record for IQ-06, Part Number: 700012245, Serial Number: 22K4B002

Deviation from Test Equipment Qualification
[bookmark: _GoBack]There were no deviations generated during the execution of the Installation Qualification, 730004135, Version 02, for the 2998 Final Test System (TE82882-8), as outlined in the Test Software Version 4.2 risk assessment, per TE82882-10JUN19-VAF.
[bookmark: _Toc309801760][bookmark: _Toc316906462][bookmark: _Toc442432510]References
The current approved versions of the following documents are referenced within this Report:
	Document Number
	Document title

	730004135
	Installation Qualification Protocol 

	735001613
	PDA Final Test System Test and Specification Procedure

	735003437
	PDA Linearity Test System Preventive Maintenance Procedure

	TE82882-10JUN19-VAF
	Test Software Version 4.2 risk assessment
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& Rev.,.[Type |Document Vs part  Description Status text Curent VeVald From |Vald To | Change No. |Long name
735001613 00 000 Acquity PDA Final Test RELEASE GENERAL 25.06.2007  30.03.2008 Eric Tumner
735001613 01 000 Acquity PDA Final Test /ABANDON Bruce Porter

B TSP 735001613 02 000 Acquity PDA Final Test RELEASE GENERAL 31.03.2008  11.05.2008 9134035  Eric Turner
9 TSP 735001613 03 000 Acquity PDA Final Test RELEASE GENERAL 12.05.2008  09.07.2008 9134937  Eric Turner
D TSP 735001613 04 000 PDA Final test RELEASE GENERAL 10.07.2008  03.12.2008 9136327  Eric Turner
E TSP 735001613 05 000 PDA Final test RELEASE GENERAL 04.12.2008  29.03.2009 9139040  Eric Turner
F TSP 735001613 06 000 PDA Final test RELEASE GENERAL 30.03.2009  21.05.2009 9140911  Eric Turner
G TSP 735001613 07 000 PDA Final test RELEASE GENERAL 22.05.2009  01.07.2010 9142852  Eric Turner
H TSP 735001613 08 000 PDA Final test RELEASE GENERAL 02.07.2010  22.11.2010 9153554  Eric Turner
) TSP 735001613 09 000 PDA Final test RELEASE GENERAL 23.11.2010  04.07.2012 9157158  Chris Green
TSP 735001613 10 000 PDA Final test /ABANDON Eric Tumner

K TSP 735001613 11 000 PDA Final test RELEASE GENERAL 05.07.2012  09.02.2014 9171410  Eric Turner
L TSP 735001613 12 000 PDA Final test RELEASE GENERAL 10.02.2014  18.02.2014 9185711  Chris Green
M TSP 735001613 13 000 PDA Final test RELEASE GENERAL 19.02.2014  12.10.2014 9186532  Eric Turner
N TSP 735001613 14 000 PDA Final test RELEASE GENERAL 13.10.2014  24.03.2015 9192039  Eric Turner
TSP 735001613 15 000 PDA Final test /ABANDON Eric Tumner

P TSP 735001613 16 000 PDA Final test RELEASE GENERAL 25.03.2015  12.07.2015 9195587  Eric Turner
R TSP 735001613 17 000 PDA Final test RELEASE GENERAL 13.07.2015  10.12.2015 9198916  Eric Turner
T TSP 735001613 18 000 PDA Final test RELEASE GENERAL 11.12.2015  28.11.2017 9202728  Tyler Model
u TSP 735001613 19 000 PDA Final test RELEASE GENERAL 29.11.2017  09.09.2018 9218596  Eric Turner
v TSP 735001613 20 000 PDA Final test RELEASE GENERAL 10.09.2018  06.12.2018 9224731  Eric Tuner
w TSP 735001613 21 000 PDA Final test RELEASE GENERAL 07.12.2018  31.12.9999 9226454  Eric Turner
T 735001613 22 000 PDA Finaltest READY Eric Tumer
TSP 735001613 23 000 PDA Final test /ABANDON Eric Tumner
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