DOCUMENT TITLE: PDA FINAL TEST

SYSTEM/ASSEMBLY NAME: PDA FINAL TEST EQUIPMENT

PART NUMBER AFFECTED See Table on Page 2

TEST EQUIPMENT NAME: PDA FINAL TEST EQUIPMENT

TEST EQUIPMENT NUMBER: TE82882, TE82969, TE83057

TEST SOFTWARE: *4.2*(TE82882)

*2.2*(TE82969)

*1.4*(TE83057)

| AUTHOR(S): eric turner | DATE: June 25, 2007 |
| --- | --- |
| REVISED BY: Matt Brule | DATE: May 20, 2019 |

# Part Numbers affected

| 186002998 | 2998 Photodiode Array Detector |
| --- | --- |
| 186015026 | ACQUITY UPLC PDA Detector |
| 186015030 | Assy., ACQUITY UPLC PDA eLambda Detector |
| 186015032 | ACQUITY UPLC PDA Detector |
| 186015033 | ACQUITY UPLC PDA eLambda Detector |
| 186015034 | ACQUITY UPC2 PDA Detector |
| 186015053 | ACQUITY PDA Detector w/TaperSlit |
| 186017003 | ACQUITY Arc 2998 PDA Detector |
| 186017013 | ACQUITY Arc Bio 2998 PDA Detector |
| 186299800 | 2998 Photodiode Array Detector (blue) |
| 279002655 | ASSY, OPTICS BENCH, 2998 |
| 700002787 | OPTICS BENCH, ACQUITY PDA |
| 700003765 | ASSY, OPTICS BENCH, 2998, CLASSIC |
| 700004035 | eLAMBDA OPTICS BENCH |
| 700005424 | OPTICS BENCH, ACQUITY PDA eLAMBDA |
| 700005427 | OPTICS BENCH, ACQUITY PDA |
| 700009014 | ASSY, OPTICS BENCH, SFC |
| 700009328 | ASSY, OPTICS BENCH, 2998 |
| 700009532 | ASSY, OPTICS BENCH, PDA |
| 700011650 | ACQUITY PDA eLAMBDA, Optics Bench |
| 700011652 | ACQUITY PDA, Optics Bench |
| 700011867 | ASSEMBLY, OPTICS BENCH, 2998 |
| 700011868 | ASSEMBLY, OPTICS BENCH, PDA TS |
| 700011869 | ASSY, OPTICS BENCH, 2998 UPC2 |
| 700012243 | Assy, Optics Bench, 2998 (On-Semi) |
| 700012245 | Optics Bench, eLambda (On-Semi) |
| 800000202 | Refurbished AcQuity UPLC PDA Detector |
| 800000244 | Refurb. 2998 Photodiode Array Detector |
| 800000245 | Ref. AcQuity UPLC PDA eLambda Detector |
| 800000304 | Ref. AcQuity PDA Detector |
| 800000305 | Ref. AcQuity PDA eLambda Det. |
| 800000397 | Ref. 2998 PDA Detector |
| 800000414 | Ref. UPC2 PDA Detector |
| 800000439 | Refurb AcQuity PDA W/Taperslit |
| 800000453 | Refurb Acq ARC 2998 PDA Det. |
| 800000476 | Refurb Arc Bio 2998 PDA Detect |
| N2971013 | Altus A-10 PDA Detector |
| N2971035 | Altus A-30 PDA Detector |

# Equipment Required

| Quantity | Part Number | Description |
| --- | --- | --- |
| 1 | TE82882 | PDA Final Test System |
| 1 | TE83057 | Detector download/configure station |
| 1 | TE82969 | PDA/2998 Linearity Test System |

# Disposables Required

| Quantity | Part Number | Description |
| --- | --- | --- |
| 1 | 176000278 | Linearity Cuvette Kit |
| 1 | 205015007 | AcQuity PDA Cuvette Holder (if AcQuity PDA or ELambda) |
| 1 | WAT042885 | Cal Solution |
| 1 | 700004155 | 2 Cuvette kit |
| 1 | 289002226 | Medium Shunt (if AcQuity PDA or eLambda) |

# reference documentation

| Document # | Description |
| --- | --- |
| 735001104 | AcQuity PDA Performance Spec |
| 735000898 | 2998 Performance Spec |
| 735001637 | eLambda Optics Performance Spec |
| 735001638 | eLambda Performance Spec |
| 735000948 | AcQuity PDA Flowcell Performance Spec |

# Definitions and Abbreviations

| Term | Definition |
| --- | --- |
| λ | Lambda |
| CPU | Central Processing Unit (i.e. “computer”) |
| FSD | Functional Specification, Design |
| L1 | Lambda 1 |
| L512 | Lambda 512 |
| LSL | Lower Spec Limit |
| MeOH | Methanol |
| PDA | Photo Diode Array |
| PRF | Performance Specification (SAP document type) |
| USL | Upper Spec Limit |
| UUT | Unit Under Test |
| SAP | Systems, Applications & Products in Data Processing (Enterprise Software) |
| PC | Personal Computer |
| LAN | Local Area Network |
| COM | Communications Port |
| AQL | Acceptable Quality Level |
| DHCP | Dynamic Host Allocation Protocol |

NOTE: For specifications and test purposes, Altus-A10, AcQuity Arc PDA, Arc Bio = 2998 Refresh and Altus-A30 = eLambda

# Test SPecifications and Traceability Matrix

| Test # | Test Specification & Condition | Acceptance criteria | Reference |
| --- | --- | --- | --- |
| 1 | Lambda 1 | ≤188.375nM(AcqPDA) ≤188nM (2998, eLambda) | PDA 735000913, sec 2.0 (see Note 3) eLambda 735001637, sec. 2.0 2998 - 735001637, sec. 2.0 |
| 2 | Lambda 512 | ≥501.25nM (AcqPDA), ≥801.5nM (2998, eLambda) | PDA 735000913, sec. 2.0 eLambda 735001637, sec. 2.0 2998 - 735001637, sec. 2.0 |
| 3 | Shunt/Flowcell energy at 5ms exposure | LSL = 80000 counts USL = 250000 counts | PSC see Note 2 |
| 4 | Linearity %RSD | ≤3.0(AcqPDA), ≤5.0 (2998, eLambda) | PDA 721007611, sec 5.15.3 eLambda 721013305, sec. 5.3.1 2998 - 721013305, sec. 5.3.2 |
| 5 | Noise | ≤7 Au (PDA, eLambda), ≤8Au (2998) | PDA 721013305, sec 5.14.1 eLambda 721013305, sec. 5.14.2 2998 - 721007611, sec. 5.15.2 |
| 6 | Drift | ≤1000 Au | PDA 721013305, sec 5.14.1 eLambda 721013305, sec. 5.14.2 2998 - 721007611, sec. 5.15.3 |

Note 1: Note 1: Optics bench temperature is reported on the test report as with specs listed with LSL as 20°C and USL as 40°C.  This is intended to automatically verify fan operation; however this is not a required test per FSD (SAP document type PRF) as operators check fan operation manually.

Note 2: Energy readings at the final test are to ensure the shunts and flow cells used are still functioning properly upon correct assembly as these energy readings are confirmed earlier in the process at the optics bench level mirror adjustment.

Note 3: On the AcQuity PDA, the performance spec allows for a maximum Lambda 1 of 188.75 so the 188.375 adheres to that spec with some additional margin.  Associated models are tested to a tighter spec than required.

NOTE: For optics bench testing, proceed to section 5

Firmware download (manual and automated) and

  - If unit under test is an AcQuity PDA, Altus A-30, or an eLambda, install medium shunt (289002226), otherwise use flowcell (289007906 or equivalent).

Remove instrument cover.

  - Verify that the battery protector has been removed.

  - NOTE: if possible, use automated test system

    - Automated

      - Install 2 pin jumper across pin 8 and 16 (software jumpers) on personality PCB to enable serial communications on blank CPU.

      - Connect UUT to power cord from power controller, serial cable to PC, and Ethernet cable to Instrument LAN in PC.

      - Make sure power switch is in the ON position. (NOTE: UUT will not turn on until power controller activates)

      - If applicable, connect leak test cable from event out connector on rear of UUT to leak detector cable connector on front of UUT.

      - If not logged into PC, log in.

      - If download-config application is not running, launch program by double clicking on the desktop icon.

      - Click Start, scan or type operator name (system will verify you are an approved operator – contact line supervisor if login error).

      - Enter serial number or scan folder serial number label with barcode scanner if available and click Enter.

      - Click Download Firmware.  Verify fan operation and air flow (out of back of UUT). Wait until complete.

      - Click Program Serial Number.

      - Click Verify.  System will check firmware version, serial number, and leak detector, if applicable

      - Collect print out, verify all tests pass, sign and place in DHR folder.  Remove jumper from CPU board.

      - Click Back To Main to test another.

NOTE: You may also power on detector with power on switch on screen.  Once detector completes power up cycle (30 sec), you can check serial number, version of firmware, and leak detector individually.

    - Manual

      - Plug UUT into available AC power.  Connect Ethernet cable to Instrument LAN card in PC.

      - Power up UUT.  Power light will turn green and lamp light will turn red.  Verify fan operation and air flow (out of back of UUT).

      - Run loader software – this should be on desktop of PC, but it can be located in the C:\Empower\Instruments\firmware directory.  Select scan for instruments box on first page, then next.

      - Program should find Boot Loader on the CPU.  Click on this and then click OK.

      - Select the proper revision of the PDA firmware (see SPECIFICATIONS section above) from the C:\Empower\Instruments\bin directory.  Click next.

      - Download should take a few minutes.  Cycle power when done. Wait for two beeps or until lamp light is blinking green. Load serial number into unit. In loader, select scan for instruments box on first page. Click next. Highlight unit, right click select (set serial number). Enter serial number of unit. OK. Click cancel to exit serial number dialog, then cancel to exit the loader program.

      - If unit is not a standard PDA (186015032), continue to section 2.

PDA (186015032) Only

      - Connect UUT to an available COM port on the PC.

      - Open Hyperterminal session with settings 38400,8,n,1.

      - Power up instrument and monitor Hyperterminal session screen.

      - After instrument displays lighting lamp, type setcalconstants 188 502 <enter>.

Verify internal fans

If instrument is not powered up, turn on power. Press control button on power controller, if applicable, to power on.

Verify that both internal fans are blowing towards the front of the unit.  The fans are under control by instrument firmware, the fans are turned on full during instrument initialization and then will slow down. this is normal.

wavelength calibration

  - If instrument is not powered up, turn on power. Press control button on power controller, if applicable, to power on.

  - Wait for UUT to complete initialization and lamp warm-up (about 4 minutes, or until lamp light is steady green).

  - Run Stop AcQuity Processes which should be on the desktop and in the C:\Empower\Instruments\bin directory.

  - If a Waters model, Run AcQuity Console which should be on the desktop and in the C:\Empower\Instruments\bin dir.

  - Click on PDA Detector from the top left of the console screen. Under configuration, view module information to verify serial number of unit under test. Before erbium calibration is performed, check to see if upper wavelength (λ512) is correct.  Under the Troubleshoot selection from the menu, select service mode, then enter 1147, for a password, OK.  Again, click on Troubleshoot, then Calibration Values.  If the unit under test is an AcQuity PDA or equiv., set the λ512 value to 510, or set to 810 if an eLambda or 2998 or equiv. is being tested. Click on Set, then Yes when warned by a dialog that you are changing calibration values.  Then click Close.

  - At this point, if this is a PDA to be tested, the lamp may have shut off due to a wavelength calibration error.  Click on the lamp icon to turn the lamp back on.  Also, reset PDA (click PDA Detector|Control|Reset PDA).

  - Under the Troubleshoot selection from the menu, select Erbium Calibration and click on the Erbium Absorbance radio button.  Once the erbium screen appears, select optimize exposure.

  - Next, click Start to continuously acquire spectrum.

  - Click on Stop, then note the Lambda 1 and Lambda 512 values.  They need to be green (L1≤188.375 and L512≥501.25 or ≥801.5 if 2998 or eLambda or equiv.).  If not, the optics bench will need to be recalibrated. If green select Save, OK, then OK again. L1 and L512 are saved into memory of unit. Close screen.

  - NOTE: 186002998, 186299800, 186015026, 186015030, 186017003, 186017013, N2971013 (and equivalent repair numbers) only - Enter Lamp Serial Number.  Select Maintain, Select Change Lamp, New Lamp, Enter serial number. (Serial number is on the Lamp cable). Select OK, Close. Close out of AcQuity Console.

  - Deactivate power controller, if applicable, by pressing control button on power controller. Unplug UUT, leaving power switch on and set up at final tester (TE82882).

Final Test (noise and drift)

  - NOTE:  If you are switching from one model to another, follow 3.1.1- 3.1.7, if not, proceed to step 3.2.

    - Open Empower by double clicking on the Empower desktop icon or find it in the All Programs list using the start button.

    - Login as system, pw: metsys.

    - Select Configure System from the menu blocks.

    - Click on Systems

    - There are up to twenty systems configured for use with the test system, PDA 1 through PDA 4, 2998 1 through 2998 4, A10 1 through A104, A30 1 through A30 4, and ELAMBDA 1 through ELAMBDA 4.  Only four of these systems can be online at a time so whatever slot on the test system you are using (1-4), the proper system needs to be on-line.  To take a system off line, right click on the system name and select take offline from the menu.  Then, right click on the system you would like to bring online and select bring online.

    - The recommended configuration is to test all of one type at a time.

    - Exit out of configuration manager by clicking on the small x in the top right corner of the window.  Exit out of Empower by clicking on the small x on the block menu window.  For extra stability, run the StopAcQuityProcesses.cmd from the desktop.

  - Connect UUT to final tester, selecting a position (1-4).  Connect the power cable, the serial cable, and the Ethernet cable from that position – they are grouped together.

  - The power switch on the UUT should be in the ON position.  The UUT will not power up, however, as it is software controlled by the tester.

  - Double click the PDA Final Test icon to start

  - Choose the loaded position and click Run.  Enter the required information and follow all instructions on the screen.  System will automatically launch Empower and run a noise and drift test.

  - Upon finishing the test, click Remove Unit, and follow instructions on screen.  A printout will print at the test station default printer.

  - Final test system will indicate if a particular UUT is to be tested for linearity (AQL testing – one of every ten, serial numbers that end in zero).

Linearity (One Unit out of 10)

  - Use PDA Linearity Test Station (TE82969).

  - Plug in UUT AC Cord, Ethernet cable, and serial cable if instrument level.

  - If unit is light guided, install cuvette holder.

  - Power up detector with MeOH cuvette (marked “0”) inside.

  - When lamp light stops blinking and is steady green, run Linearity test program (if not already running) and press START and fill in all required information and select appropriate report mode (Production for standard production and Repair for service).  Note: make sure “Launch Console” switch on computer screen is in the ON position (default).

  - Wait 3-4 seconds (do not click start again), and the program will launch the console application.

  - Click on PDA – top left corner of screen.  Graphs should appear on screen with information on detector.

  - Click on Troubleshoot|Service Mode

  - Enter 1147 as password, OK.

  - Click Maintain|Read Energy

  - Click the Continuous Read button to select it. Set wavelength to 240 nm.  Type in an exposure of 12.00, and then click Read.  Note the peak height at around 240nM. Stop (Click the Read button again which is now labeled Stop) and start (Click Read) reading energy, adjusting exposure time until peak between 170,000 and 220,000 counts.

  - After determining the correct exposure, click Stop.  Click the Continuous Read button to select it.

  - Click on the close shutter box.  Click Read once to take a scan. Click Read, click Capture. Repeat Read-Capture two more times for a total of 3.  Make sure screen shows Captures:3.

  - Press Save.  A dialog box will open, looking for a filename.  Name File C:\PDA Linearity\TestData\darkxxx.txt where xxx is the last 3 numbers in the serial number (before the factory specific letter).  If an optics bench, use the entire optics bench serial number as xxx.  Click Save.

  - Open the shutter.  Make sure exposure is the same as the last capture.

  - Click Read once to take a scan. Click Read, click Capture. Repeat Read-Capture two more times for a total of 3.  Make sure screen shows Captures 3 then stop.

  - Press Save.  Name File C:\Program Files\PDA Linearity\TestData\meohxxx.txt where xxx is the last 3 numbers in the serial number (before the factory specific letter). If an optics bench, use the entire optics bench serial number as xxx.  Click Save.

  - Repeat Steps 17 and 18 for 5mg, 10mg, 15mg, 20mg, and 25mg cuvettes, also using a 30mg cuvette for extended wavelength detectors, naming the text file 5mgxxx.txt, 10mgxxx.txt etc.  If an optics bench, use the entire optics bench serial number as xxx.  Exit Read Energy.

  - Exit Console.  Graph will show and print linearity points and %RSD.  If %RSD is over 3.0(std) or 5.0(extended wavelength) retest.  If linearity does not go below 3.0(std) or 5.0(extended wavelength) within three retests, contact engineering.  Insert passing printout into file.

spare Optics Bench testing

  - Install bench into a new or tested chassis which has appropriate firmware.  Follow test procedure from Step 2.1, using the download station to run the console application (to get erbium cal data), and then run final test, section 3.

  - The chassis will be left with the last serial number tested in its static RAM, but when the chassis is used for a shippable unit, the correct serial number will be programmed in.

troubleshooting

  - If the system has trouble launching console or Empower, run Waters DHCP Server Configuration.exe and verify there are less than 255 IP addresses, serial numbers in list.  If not, delete them all.  If you cannot locate the program do a search of *DHCP* with the XP search function.

  - You may also need to run “Stop AcQuity Processes.cmd“ if you are having connection problems or, if all else fails, re-boot system.

  - Verify the proper systems are on-line (4.1.1-4.1.7).

  - To verify a connection to a single system

    - Plug AC into a regular power outlet and power up

    - Open HyperTerminal and connect to the appropriate serial port (COM15 for Station 1, COM16 for station 2 etc.) with 38400 baud, 8, n,1.

    - Change the serial number of the unit to 1-4 for AcqPda, position 1-4, 5-8 for 2998, position 1-4, or 9-12 for eLambda, position 1-4, 13-16 for A10, 17-20 for A30.  eg.  2998 in position 2 would be setserialnumber 6 <enter>.

    - Type setipflags N F <enter>.

    - Type setipaddress 192 168 0 x <enter> where x = 5-8 for AcqPda 1-4, x=9-12 for 2998 1-4, and x=13-16 for eLambda 1-4.  So, for an eLambda in position 2, you would type setipaddress 192 168 0 14 <enter>.

    - Power cycle UUT.

    - Plug Ethernet cable directly into switcher, or card in PC.

    - Launch empower and login as system pw:metsys.  Click Configure System.

    - Select Empower Nodes and right click on the empower node on the right.  Select properties from the drop-down list.

    - Select Instruments and check to see if the instrument in question is “OK”.  You may run scan instruments to see if it helps.  If not, call TE.  Sometimes these steps “awaken” Empower to now recognize the UUT.

  - To troubleshoot Lost Comm errors:

    - In the list of instruments, there should only be up to 20 systems (Address ACQ-PDA#1 through #4, and #9 through #12, W2998#5 through #8, A10#13-16, A30#17-20).  If there are any other systems listed with typical serial numbers (A09UPD011M) or ‘serial number uninitialized’, click remove.

    - Click on the Configure DHCP tab.  Click on the Configure DHCP button.  If there are any instruments listed in the Waters DHCP server configuration list, highlight them and remove them.  Click OK.

    - Click OK, then use the X in the top corner of the Configuration Manager window and the then the Empower Pro window to exit.

    - Put any removed cabling back and try running application again.

    - If problems persist, reboot system.

    - To minimize Lost Comm errors, only run the same model detectors at the same time.

    - If condition still persists, exit application and click on Stop AcQuity Processes.

    - Click on Start|Run and type in net stop "Waters DHCP Server".

    - After service stops, delete C:\Empower\ Instruments\DHCP.xml.

    - Restart test program.

    - If you are still facing connection issues, try the following:

      - Start the test program and start the test on any position with a unit attached.

      - When the test software displays "Powering Up", click on the X in the top right corner of the test software window to close the program.  This will leave the Ethernet connected and the power on.

      - Restart PC.

      - Try running Loader.exe (C:\Empower\Instruments\Firmware folder).  Check the scan for instruments box and click OK.  Sometimes this establishes a connection which you will see in the information box.  Click cancel and exit loader.

      - To verify connection at this point, launch Empower, user tester, no password.

      - Click on Configure System.

      - Choose Empower Nodes from tree.

      - Right click on the node name and select Properties…

      - Select the Instrument tab and look to see if the instrument connected shows Yes in the OK? column.

      - If yes, close Empower and restart the test program, which will power off the instrument.  You must power up the instrument again and reset the serial number before testing.

# END OF DATA