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Equipment Required

	Quantity
	Part Number
	Description

	1
	TE83026-1
	Acquity CMA Test system

	
	
	


Disposables Required

	Quantity
	Part Number
	Description

	
	
	

	
	
	


reference documentation

	Document Number
	Description

	WAT046538
	Electrical Safety Procedure

	
	


SPecifications

All LEDs are verified

Fans are checked

Interfaces Tested:

· 4 Pre-Heater Interfaces

· 2 I-Button Interfaces
· Leak detector

Heater Cooler functionality is verified by heating to 90 degrees C, Cooling to 4 degrees C, 
and then returning to 25 degrees C, in both column compartments.
This procedure assumes that the operator has a basic knowledge of personal computer operation in a Microsoft Windows environment, running Water’s Empower Software.

All figures in this procedure are for reference only and may differ from actual screens etc.

To shut down a test workstation:

Close all programs.

Click on Start.

Shut Down.

Select Shut Down the Computer?

Select Yes.

Failure to follow these instructions may result in workstation damage.

Test System Setup
Setup TE83026 PC with monitor / keyboard / mouse / barcode scanner (all USB).

Connect relay controller to a USB port.

Connect the power pack to the relay module

Connect the AC power controller to COM2 of the PC.

Connect The AC power controller to appropriate power source.

Connect the PC to the network.

Connect COM1 to the Know good CMA

Connect the Known good CMA to power with line cord from power Controller.
Power up the PC

Login: tester / password: tester.

Launch the TE83026-AUX test program from the desktop of Start menu.

Test System Controls
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Figure 1 (Reference Only)
UUT connections:

AUX Cable, line cord from direct power source, leak detector cable.
1. UUT Test Setup

1.1. Connect the power cord from the direct power source to the UUT.
1.2. Connect the AUX cable to the Upper Aux port an the known good CMA..
1.3. Connect the 3 pin leak detector cable from the test system to the UUT.

2. Electrical Safety Test

2.1. Perform per procedure WAT046538.
3. Test

3.1. Press the Start Button under Automated test with the mouse.

3.2. Follow the on-screen instructions and supply the requested data, by scanning with the barcode scanner or typing.  If something is entered that the tester does not support or recognize, the system will prompt for the input to be repeated.

The test will proceed sequentially through each of the tests listed in the “Test Status” list.  Pertinent test information will be displayed.  Follow all instructions to complete testing.  When the test is complete a Test Record will be generated, data will be written to the Sigmaquest data collection system.  Either pass or fail will be displayed (see Figure 1 for test screen layout).

3.3. When the test is complete:

3.3.1. Push the UUT power button on CM-A (Known good unit) to power it off.
3.3.2. Disconnect all cables.

3.3.3. Remove  I-buttons.
3.3.4. Remove all Pre-heaters.

3.4. Verify that the UUT passed, mark the test record accordingly, and sign and date the test record.
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