| Test Script # |  |  | IQ-01 |  |  |  |  | IQ Test Name: |  |  | Associated Documentation Verification |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Equipment # |  |  | TE or Val # |  |  |  |  | Equipment Name: |  |  | Enter the equipment name & serial number/ID, if applicable |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Installation Site: |  |  | Enter the equipment installation site, i.e. Milford Production Clean Room |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Purpose: To verify that supporting documentation is available prior to the completion of qualification. Test Execution: - List all applicable documents (Preventive Maintenance, Calibration, Operating Procedures, Test Specification Procedures, Manufacturing Procedures, etc.) under the Document type header, then complete the table below. Rows may be added/removed as needed. - All applicable documents must have a status of Ready, Pre-Rel-Prod, or Release General in SAP. If applicable documents are unavailable in SAP state why in the comments section. Objective Evidence: Completed IQ-01 Test Script Data Sheet SAP screenshots of below available document(s) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Document Type |  |  |  |  | Title, Doc. #, Version/Revision # |  |  |  |  |  |  |  | Status in SAP |  |  |  |  |  |  |  | Initial & Date |  |  |  |
| Enter the Document Type from SAP, i.e. TSP |  |  |  |  | Enter the Title Document number, Version/Revision numbers |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Enter the Document Type from SAP, i.e. Cal |  |  |  |  | Enter the Title Document number, Version/Revision numbers |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Comments: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Overall Test Conclusion |  |  |  |  |  |  |  |  |  |  | Document No.: ______________________. Iteration No.: ________. Attachment: _______ Page: ___ of ___. |
| Pass Fail |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Test Script # |  |  | IQ -02 |  |  |  | IQ Test Name: |  |  |  | Site Verification |  |  |  |  |  |  |  |  |  |  |  |  |
| Equipment # |  |  | TE or Val # |  |  |  | Equipment Name: |  |  |  | Enter the equipment name & serial number/ID, if applicable |  |  |  |  |  |  |  |  |  |  |  |  |
| Installation Site: |  |  | Enter the equipment installation site, i.e. Milford Production Clean Room |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Purpose: To verify the location in which the test equipment is to be installed is suitable for use. Test Execution: - List additional applicable environmental/input parameters in the Parameters column. Rows may be added as needed. - Using the applicable equipment documentation, define the specification(s) in the Acceptance Criteria column, as appropriate. - Verify and/or measure the applicable parameters and complete the table below. Do not delete any of the listed defined parameters. Enter N/A for any parameter or cell that is not used or applicable. The reason for why the parameter is not used or applicable shall be explained within the corresponding Installation Qualification Report. Objective Evidence: Completed IQ-02 Test Script Data Sheet Photo(s) of the work area, as applicable The source of each Acceptance Criterion shall be referenced in the IQR if stored in SAP. If not in SAP objective evidence specifying the Acceptance Criterion shall be provided in or as an attachment to the IQR. Copies of calibration certificates for measurement tools, if applicable |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Parameters |  |  |  | Acceptance Criteria |  |  |  |  | As Found |  |  |  |  |  |  |  | Pass/Fail |  |  |  |  | Initial & Date |  |
| Ambient Temperature |  |  |  | 4 – 40 °C |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Relative Humidity |  |  |  | 10 – 90 % Non-condensing |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Workspace |  |  |  | Adequate space for testing and service |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Test Station Bench |  |  |  | Large and strong enough to support all equipment |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Electrical Supply |  |  |  | 100 – 120 VAC 47 – 63 Hz |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Electrical Outlets |  |  |  | ≥ |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Network Available |  |  |  | ≥ |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Comments: Temperature and humidity specified in Electrical supply specified in |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Overall Test Conclusion |  |  |  |  |  |  |  |  |  |  | Document No.: ______________________. Iteration No.: ________. Attachment: _______ Page: ___ of ___. |
| Pass Fail |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Test Script # |  | IQ-03 |  |  |  |  | IQ Test Name: |  |  | Instrumentation Calibration Verification |  |  |  |  |  |  |  |  |  |  |  |  |
| Equipment # |  | TE or Val # |  |  |  |  | Equipment Name: |  |  | Enter the equipment name & serial number/ID, if applicable |  |  |  |  |  |  |  |  |  |  |  |  |
| Installation Site: |  | Enter the equipment installation site, i.e. Milford Production Clean Room |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Purpose: To verify that all critical instruments are in the calibration database and that a calibration and/or a preventive maintenance has been performed, as appropriate. Test Execution: List all critical instruments/components and/or sub-systems that require calibration and/or preventive maintenance in the Description column. Rows may be added as needed. Verify that all listed critical instruments/components have had a preventive maintenance and/or calibration performed in accordance with the PM and/or Calibration procedure(s) or by an outside source, and are listed within the associated calibration database, then complete the table below. Objective Evidence: Completed IQ-03 Test Script Data Sheet Copies of calibration and/or maintenance certificates Screen captures from the associated Calibration database |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Description |  |  |  |  |  | Asset # |  |  |  | Cal/PM Dates |  |  |  |  |  |  |  | Initial & Date |  |  |  |  |
| Enter the equipment or tooling nomenclature |  |  |  |  |  |  |  |  |  | Last Cal/PM: Cal/PM Due: |  |  |  |  |  |  |  |  |  |  |  |  |
| Comments: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Overall Test Conclusion |  |  |  |  |  |  |  |  |  | Document No.: ______________________. Iteration No.: ________. Attachment: _______ Page: ___ of ___. |
| Pass Fail |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Test Script # |  |  | IQ-04 |  |  |  | IQ Test Name: |  |  |  | Test Software and Instrument Firmware Version |  |  |  |  |  |  |  |  |  |  |  |  |
| Equipment # |  |  | TE or Val # |  |  |  | Equipment Name: |  |  |  | Enter the equipment name & serial number/ID, if applicable |  |  |  |  |  |  |  |  |  |  |  |  |
| Installation Site: |  |  | Enter the equipment installation site, i.e. Milford Production Clean Room |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Purpose: To verify that the installed software/firmware is correct per the applicable design and/or operational document(s). Test Execution: List all relevant Equipment software and/or instrument firmware in the Description column. Specify the associated part number for the instrument firmware, if applicable. Rows may be added/removed as needed. List the associated required equipment software and/or critical instrument/internet connection sharing firmware version(s), as defined in the applicable design/operational document(s), in the Acceptance Criteria column. Verify that the software version displayed on the Operator Interface of the program and/or the firmware loaded on the instruments/PC match the acceptance criteria identically, if applicable, then complete the table below. Objective Evidence: Completed IQ-04 Test Script Data Sheet Identify and/or provide a screen capture/print out/copy of the applicable design/operational document that defines the required software/firmware version(s) within or as an attachment to the associated Installation Qualification Report. Screen capture and/or test record displaying/indicating the software and/or firmware version(s), as applicable |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Description |  |  |  |  |  | Acceptance Criteria |  |  |  |  | As Found |  |  |  | Pass/Fail |  |  |  |  | Initial & Date |  |  |  |
| Test Equipment Software |  |  |  |  |  | Version: 1.0 |  |  |  |  | Version: |  |  |  |  |  |  |  |  |  |  |  |  |
| Instrument Firmware P/N: 186017013 |  |  |  |  |  | Version: 1.65.6205 |  |  |  |  | Version: |  |  |  |  |  |  |  |  |  |  |  |  |
| Comments: Test software version specified in |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Overall Test Conclusion |  |  |  |  |  |  |  |  |  |  | Document No.: ______________________. Iteration No.: ________. Attachment: _______ Page: ___ of ___. |
| Pass Fail |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Test Script # | IQ-05 |  |  |  |  |  | IQ Test Name: |  |  |  | General Safety and Alarms |  |  |  |  |  |  |  |  |  |  |  |  |
| Equipment # | TE or Val # |  |  |  |  |  | Equipment Name: |  |  |  | Enter the equipment name & serial number/ID, if applicable |  |  |  |  |  |  |  |  |  |  |  |  |
| Installation Site: |  | Enter the equipment installation site, i.e. Milford Production Clean Room |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Purpose: To verify that the equipment can operate in a safe manner, within its designated location. Test Execution: List any additional safety related conditions or alarms in the Safety Inspection column, such as: guarding, interlocks, thermal or mechanical safety, chemical, secured gas tanks, restart or power loss, emergency stop, environmental exhaust, electrical safety, noise, safety labels, etc. Rows may be added as needed. EH&S shall be consulted as necessary. List the associated safety requirement(s)/specification(s) in the Acceptance Criteria column. Perform an inspection/test of the equipment to assess its safety related conditions, then complete the table below. Objective Evidence: Completed IQ-05 Test Script Data Sheet Photo(s) of the secured glass bottles, tanks, etc., if applicable The source of any applicable Acceptance Criterion shall be summarized in the IQR if stored in SAP. If the source is not stored in SAP, at a minimum the page specifying the applicable Acceptance Criterion shall be provided as an attachment. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Safety Inspection |  |  |  |  |  |  | Acceptance Criteria |  |  |  |  |  |  |  |  | Pass/Fail |  |  | Initial & Date |  |  |  |  |
| Mechanical Safety/ Pinch Points/Sharp Edges |  |  |  |  |  |  | None Found (or mitigated appropriately) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Glass Bottles |  |  |  |  |  |  | Secure |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Waste Tanks |  |  |  |  |  |  | Secure and located below the system |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Emergency Stop |  |  |  |  |  |  | System ceases operation upon depressing the Emergency Stop button |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Guarding |  |  |  |  |  |  | Shielding in place to prevent hazard during operation |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Comments: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Overall Test Conclusion |  |  |  |  |  |  |  |  |  |  |  | Document No.: ______________________. Iteration No.: ________. Attachment: _______ Page: ___ of ___. |
| Pass Fail |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Test Script # | IQ-06 |  |  |  |  |  | IQ Test Name: |  |  |  | Functional Verification |  |  |  |  |  |  |  |  |  |  |  |  |
| Equipment # | TE or Val # |  |  |  |  |  | Equipment Name: |  |  |  | Enter the equipment name & serial number/ID, if applicable |  |  |  |  |  |  |  |  |  |  |  |  |
| Installation Site: |  | Enter the equipment installation site, i.e. Milford Production Clean Room |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Purpose: To verify that the system operates as expected. Test Execution: Select the UUT(s) or tests to be performed as defined within the Installation Qualification Section of 730001835. List the Part Number and Serial Number, if applicable, in the UUT Part Number/Serial Number column. Define the appropriate Acceptance Criteria within the Acceptance Criteria column. Perform a test of the UUT(s), verify the Acceptance Criteria, then complete the table below. If an associated Test Record is being used as objective evidence the Header/Footer information should be verified for accuracy. Objective Evidence: Completed IQ-06 Test Script Data Sheet Copy of the Test Record(s) for the UUT(s), as applicable The UUT(s) selected for this Test Script and the reasoning for why it/they were selected shall be summarized within the associated Installation Qualification Report. If this Acceptance Criteria is other than “All required CTQ Specification measurements result in a Pass” then it shall be summarized within the associated VAF, Process Validation Plan, and/or Installation Qualification Report. For systems that download firmware to product the BOM showing the firmware requirement shall be provided within the IQR. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| UUT Part Number/Serial Number |  |  |  |  |  |  | Acceptance Criteria |  |  |  |  |  |  | Pass/Fail |  |  |  | Initial & Date |  |  |  |  |  |
| Part Number: Serial Number: |  |  |  |  |  |  | All required CTQ Specification measurements result in a Pass |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  | UUT Firmware Version: 1.40.45 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Comments: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Overall Test Conclusion |  |  |  |  |  |  |  |  |  |  | Document No.: ______________________. Iteration No.: ________. Attachment: _______ Page: ___ of ___. |
| Pass Fail |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

Appendix A:  TRAINING RECORD

| Document # |  | Document Title |  | Version |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  |  | Revision |  |

Read and Understood

By signing you acknowledge that you have reviewed, understand, and agree to comply with the policies and practices of the procedure defined above.

| Printed Name | Signature | Initials | Date |
| --- | --- | --- | --- |

|  | Document No.: ______________________. Iteration No.: ________. Attachment: _______ Page: ___ of ___. |
| --- | --- |