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	[bookmark: _Hlk14685499]


Test Script #     
	IQ-01
	IQ Test Name:     
	Associated Documentation Verification

	[bookmark: _Hlk14180957]Equipment #     
	TE82882-13
	[bookmark: _Hlk14181388]Equipment Name:     
	PDA Final Test System

	Installation Site:     
	Flex, Singapore

	Purpose: 
To verify that supporting documentation is available prior to the completion of qualification.
Test Execution:
· List all applicable documents (Preventive Maintenance, Calibration, Operating Procedures, Test Specification Procedures, Manufacturing Procedures, etc.) under the Document type header, then complete the table below. Rows may be added/removed as needed. 
· [bookmark: _GoBack]All applicable documents must have a status of Ready, Pre-Rel-Prod, or Release General in SAP. If applicable documents are unavailable in SAP state, why in the comments section. 
Objective Evidence:
Completed IQ-01 Test Script Data Sheet 
SAP screenshots of below available document(s) 

	Document Type
	Title, Doc. #, Version/Revision #
	Status in SAP
	Initial & Date

	TSP
	PDA Final Test System Test and Specification Procedure, 735001613, Version 25 
	Ready
	

	Comments:


















	Overall Test Conclusion
		
	Document No.: ______________________

	
	Iteration No.: ________

	
	Attachment: _______ Page: ___ of ___




	 Pass			 Fail
	

	



Test Script #
	IQ-06
	IQ Test Name:
	Functional Verification

	Equipment #
	TE82882-13	Equipment Name:
	PDA Final Test System
	Installation Site:
	Flex, Singapore
	Purpose:
To verify that the system operates as expected.
Test Execution:
Select the UUT(s) or tests to be performed as defined within the Installation Qualification Section of 730001835.
List the Part Number and Serial Number, if applicable, in the UUT Part Number/Serial Number column.
Define the appropriate Acceptance Criteria within the Acceptance Criteria column.
Perform a test of the UUT(s), verify the Acceptance Criteria, then complete the table below. If an associated Test Record is being used as objective evidence the Header/Footer information should be verified for accuracy.
Objective Evidence:
Completed IQ-06 Test Script Data Sheet
Copy of the Test Record(s) for the UUT(s), as applicable
The UUT(s) selected for this Test Script and the reasoning for why it/they were selected shall be summarized within the associated Installation Qualification Report.
If this Acceptance Criteria is other than “All required CTQ Specification measurements result in a Pass” then it shall be summarized within the associated VAF, Process Validation Plan, and/or Installation Qualification Report.  
For systems that download firmware to product the BOM showing the firmware requirement shall be provided within the IQR.

	UUT Part Number/Serial Number
	Acceptance Criteria
	Pass/Fail
	Initial & Date

	  Part Number:    186018007
  Serial Number:  
	 All required CTQ Specification measurements result in a Pass 
	Pass
	

	Comments:  


	Overall Test Conclusion
		
	Document No.: ______________________

	
	Iteration No.: ________

	
	Attachment: _______ Page: ___ of ___




	 Pass			 Fail
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