| Test Script # | IQ-01 |  | IQ Test Name: | Associated Documentation Verification |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Equipment # | TE82943-12 |  | Equipment Name: | ACQUITY SDS Automated Test System |  |  |
| Installation Site: | Flex, Singapore |  |  |  |  |  |
| Purpose: To verify that supporting documentation is available prior to the completion of qualification. Test Execution: - List all applicable documents (Preventive Maintenance, Calibration, Operating Procedures, Test Specification Procedures, Manufacturing Procedures, etc.) under the Document type header, then complete the table below. Rows may be added/removed as needed. - All applicable documents must have a status of Ready, Pre-Rel-Prod, or Release General in SAP. If applicable documents are unavailable in SAP state why in the comments section. Objective Evidence: Completed IQ-01 Test Script Data Sheet SAP screenshots of below available document(s) |  |  |  |  |  |  |
| Document Type |  | Title, Doc. #, Version/Revision # |  |  | Status in SAP | Initial & Date |
| TSP |  | acquity BSM PUMP AUTOMATED FINAL TEST test procedure, 735002010 Revision P, VERSION 13 |  |  | RELEASE GENERAL | HZ 22 OCT 2020 |
| CAL |  | calIBRATION/MAINTENANCE FOR te82943, ACQUITY SDS AUTOMATED TESTER, 735001828, rEVISION E, Version 04 |  |  | RELEASE GENERAL | HZ 22 OCT 2020 |
| Comments: |  |  |  |  |  |  |
| Overall Test Conclusion |  |  |  | Document No.: ______________________. Iteration No.: ________. Attachment: _______ Page: ___ of ___. |
| ☒ Pass Fail |  |  |  |  |  |  |

| Test Script # | IQ-04 |  | IQ Test Name: | Test Software and Instrument Firmware Version |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Equipment # | TE82943-12 |  | Equipment Name: | ACQUITY SDS Automated Test System |  |  |
| Installation Site: | Flex, Singapore |  |  |  |  |  |
| Purpose: To verify that the installed software/firmware is correct per the applicable design and/or operational document(s). Test Execution: List all relevant Equipment software and/or instrument firmware in the Description column. Specify the associated part number for the instrument firmware, if applicable. Rows may be added/removed as needed. List the associated required equipment software and/or critical instrument/internet connection sharing firmware version(s), as defined in the applicable design/operational document(s), in the Acceptance Criteria column. Verify that the software version displayed on the Operator Interface of the program and/or the firmware loaded on the instruments/PC match the acceptance criteria identically, if applicable, then complete the table below. Objective Evidence: Completed IQ-04 Test Script Data Sheet Identify and/or provide a screen capture/print out/copy of the applicable design/operational document that defines the required software/firmware version(s) within or as an attachment to the associated Installation Qualification Report. Screen capture and/or test record displaying/indicating the software and/or firmware version(s), as applicable |  |  |  |  |  |  |
| Description |  | Acceptance Criteria |  | As Found | Pass/Fail | Initial & Date |
| ACQUITY SDS Final Test Station System Software |  | Version: 2.4 |  | Version: 2.4 | PASS | HZ 22 OCT 2020 |
| Comments: |  |  |  |  |  |  |
| Overall Test Conclusion |  |  |  | Document No.: ______________________. Iteration No.: ________. Attachment: _______ Page: ___ of ___. |
| ☒ Pass Fail |  |  |  |  |  |  |

| Test Script # | IQ-06 | IQ Test Name: |  | Functional Verification |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Equipment # | TE82943-12 | Equipment Name: |  | ACQUITY SDS Automated Test System |  |  |  |
| Installation Site: | Flex, Singapore |  |  |  |  |  |  |
| Purpose: To verify that the system operates as expected. Test Execution: Select the UUT(s) or tests to be performed as defined within the Installation Qualification Section of 730001835. List the Part Number and Serial Number, if applicable, in the UUT Part Number/Serial Number column. Define the appropriate Acceptance Criteria within the Acceptance Criteria column. Perform a test of the UUT(s), verify the Acceptance Criteria, then complete the table below. If an associated Test Record is being used as objective evidence the Header/Footer information should be verified for accuracy. Objective Evidence: Completed IQ-06 Test Script Data Sheet Copy of the Test Record(s) for the UUT(s), as applicable The UUT(s) selected for this Test Script and the reasoning for why it/they were selected shall be summarized within the associated Installation Qualification Report. If this Acceptance Criteria is other than “All required CTQ Specification measurements result in a Pass” then it shall be summarized within the associated VAF, Process Validation Plan, and/or Installation Qualification Report. For systems that download firmware to product the BOM showing the firmware requirement shall be provided within the IQR. |  |  |  |  |  |  |  |
| UUT Part Number/Serial Number |  |  | Acceptance Criteria |  |  | Pass/Fail | Initial & Date |
| Part Number: 186015061 Serial Number: J20BBS001A |  |  | All required CTQ Specification measurements result in a Pass |  |  | PASS | HZ 22 OCT 2020 |
| Comments: |  |  |  |  |  |  |  |
| Overall Test Conclusion |  |  |  |  | Document No.: ______________________. Iteration No.: ________. Attachment: _______ Page: ___ of ___. |
| ☒ Pass Fail |  |  |  |  |  |  |  |