# installation qualification REPORT

For the Metering Syringe Valve Leak Test System (TE0084040-1)

PREPARED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Tyler Modelski | Test Engineer II |  |  |

REVIEWED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Emma Pierce | Test Validation Engineer |  |  |

APPROVED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Nick Grow | Quality Internal Audit Coordinator |  |  |

The “Prepared By” signature denotes that this document accurately describes the qualification activities.

The “Reviewed By” signature[s] denotes agreement with the description of the qualification activities.

The “Approved By” signature denotes that this document has been reviewed and approved for overall quality and completeness.

REVISION HISTORY

| Version | Author | Summary of change |
| --- | --- | --- |
| 00 | Tyler Modelski | First document for approval |

# Purpose

The Metering Syringe Valve Leak Test System (TE0084040-1) is located in the White Room in Milford, MA. It is used to perform leak testing of the following part numbers: 289010714 and 289012763. These part numbers are functionally equivalent, the only difference being material (stainless steel versus titanium). They are tested by the system identically; therefore, the IQ was only performed using the 289010714 part number. Testing is achieved by running a leak test for each port to port configuration of the valve:

- Port 2 to 3: Inject Valve to Purge/Wash

- Port 1 to 3: Syringe to Purge/Wash

- Port 1 to 2: Syringe to Inject Valve.

The test system is comprised of a Micro BSM, Bronkhurst Flow Meter, and two selector valves alongside a test computer installed with custom software written in National Instruments LabVIEW 2017. The Micro BSM is used to deliver flow and to set pressure within the fluidics of the test system. The selector valves are used to set the flow path through the correct UUT port and flow to waste/dead end the system. The leak rate is measured using the flow meter. A test fixture (999045028) is used to move the valve to each port to port fluid path. This fixture is manually operated by the test operator to position the valve in the correct path.

# Scope of Report

This Installation Qualification Report details the qualification activities undertaken for the Metering Syringe Valve Leak Test System (TE0084040-1) based on the requirements of the executed protocol, Installation Qualification Procedure 730004135 Version 00.

This Installation Qualification Report defines the documents used to qualify the Metering Syringe Valve Leak Test System (TE0084040-1).

# Summary of Results - Statement of Fitness for Intended Use

The Metering Syringe Valve Leak Test System (TE0084040-1) was tested in accordance Installation Qualification Procedure 730004135 Version 00.  All test scripts were performed and all objective evidence requirements have been met.

Waters now has the necessary documentation in place to ensure that the Metering Syringe Valve Leak Test System (TE0084040-1) complies with Waters’ requirements, the latest relevant quality system regulations, and cGMP practices.

Following the approval of this Installation Qualification Report, the Metering Syringe Valve Leak Test System (TE0084040-1) can be considered fully qualified and fit for intended use.

# Glossary

The following definitions are used in this document:

| Term | Explanation |
| --- | --- |
| µL/min | Microliters Per Minute |
| BSM | Binary Solvent Manager |
| Cal/PM | Calibration/Preventive Maintenance |
| cGMP | Current Good Manufacturing Practice |
| EDV | Equipment Design Verification |
| IQ | Installation Qualification |
| MSV | Metering Syringe Valve |
| P/N | Part Number |
| SAP | Enterprise Resource Planning (ERP) software |
| TMV | Test Method Validation |
| TSP | Test Specification and Procedure |
| UI | User Interface |
| UUT | Unit Under Test |

# Training

The executor of the Installation Qualification Procedure 730004135 Version 00 protocol was trained. Refer to Attachment #1 for a screenshot of the training database. The executor of script IQ-06 was also the author of the TSP; therefore, no additional training was required.

# Installation Qualification Results

The table below provides a summary of the results of the scripts executed as part of the Installation Qualification Procedure 730004135 Version 00. The objective evidence showing the results for each test step can be found in the attachments specified in the table.

| Script or Section | Pass/Fail | Objective Evidence | Deviations |
| --- | --- | --- | --- |
| IQ-01 | Pass | Attachments #2 and #3 | None |
| IQ-02 | Pass | Attachments #4, #5, #6 | None |
| IQ-03 | Pass | Attachments #7 and #8 | None |
| IQ-04 | Pass | Attachments #9 and #10 | None |
| IQ-05 | Pass | Attachments #11 and #5 | None |
| IQ-06 | Pass | Attachments #12 and #13 | None |

# Attachments

Attachment #1: Training to 73004135 Installation Qualification Procedure – Screenshot of Database

Attachment #2: Completed Test Script Data Sheet IQ-01: Associated Documentation

Attachment #3: Objective Evidence IQ-01: SAP Screenshots of TSP 735004416 and CAL/PM 735004437

Attachment #4: Completed Test Script Data Sheet IQ-02: Site Verification

Attachment #5: Objective Evidence IQ-02 and IQ-05: Photo of Test Setup

Attachment #6: Objective Evidence IQ-02: Calibration Certificate Gauges/Meters

Attachment #7: Completed Test Script Data Sheet IQ-03: Instrument Calibration Verification

Attachment #8: Objective Evidence IQ-03: Screenshots of Calibration / Preventative Maintenance Database

Attachment #9: Completed Test Script Data Sheet IQ-04: Test Software and Instrument Firmware Version

Attachment #10: Objective Evidence IQ-04: Screenshots of Test Software Version and Micro BSM Firmware Version

Attachment #11: Completed Test Script Data Sheet IQ-05: General Safety and Alarms

Attachment #12: Completed Test Script Data Sheet IQ-06: Functional Verification

Attachment #13 Objective Evidence IQ-06: Screenshot of Test Software UI Displaying Passing UUT Result

# Deviation from Installation Qualification

There were no deviations generated during the execution of Installation Qualification Procedure 730004135 Version 00.

# References

The current approved versions of the following documents are referenced within this document:

| Document Number | Document title |
| --- | --- |
| 730004135 | Installation Qualification Procedure |
| 735004416 | Test Spec and Procedure, MSV Leak Test |
| 735004437 | Cal/PM TE0084040 MSV Leak Test System |
| 735004417 | Metering Syringe Valve Subassembly Specification |
| TE0084040-TMVR Iteration 00 SAP: TE0084040QX Version 00) | Test Method Validation Report MSV Leak Test System TE0084040 |
| TE0084040-EDVR Iteration 00 SAP: TE0084040QX Version 00) | Equipment Design Verification MSV Leak Test System TE0084040 |