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IQ-02 TEST SCRIP DATA SHEET: 

AMBIENT TEMPERATURE: 21 ºC to 22 ºC
Verified using 
THERMOHYGROGRAPH
Serial Number : 5650
EQID : 0418EQID : 0418
Calibration Due Date: 07-Jan-2021
[image: ]







RELATIVE HUMIDITY: 51 % to 55 %
Verified using 
THERMOHYGROGRAPH
Serial Number : 5650
EQID : 0418
Calibration Due Date: 07-Jan-2021
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ELECTRICAL SUPPLY : 
Verified using : 
FLUKE 110 TRUE RMS MULTIMETER
Serial Number: 85730619
EQID: 0245
Calibration Due Date: 16-Dec-2020
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Minimum 4 dedicated outlets

ELECTRICAL OUTLET 
Computer Work
Station (AC Line CPU)
LCD Monitor (AC Line_

TE82526 Pump 
Tray Tester (AC Line)
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AND HR 200 Electronics Balance  (AC Line)
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NETWORK AVAILABLE: Minimum of 1 dedicated network
Blue Sticker : Dedicated network for Pump Tray Area
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WS-9 is a dedicated network












WORK SPACE: The system has sufficient room for the test station, UUT and service
TE82526 Pump 
Tray Tester
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TEST STATION BENCH :   The test station bench is large and strong enough to support the station                                                      
[image: ]Test Station Bench






IQ-04: Test Software and Instrument Firmware Version
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IQ-04: Test Software and Instrument Firmware Version
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IQ-05 Test Script Data Sheet:

Station does not have Pinch Points/ sharp edges
Glass Bottle is secured by plastic containerGlass Bottles secured inside the container

[bookmark: _GoBack][image: ]Waste Bottles secured inside the container

                Wastes Bottle / Glass  secured inside the container
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THE SCIENCE OF WHAT'S POSSIBLES Title: individual Training Form

Document No: 730002657 Version No: 00 Page 1 0of 1

Training Type: | Read and Understood

Name: | Luo HuaZhen Em;l)ll;yee 618078

Date of Training | 19SEP2019

Comments: | N/A

Document Number Title Rev. #

( 730004135 Installation Qualification Protocol 03

% (¢

O Gontoy_ Tt e 19502013

Employee’s Signature Title Date
W (95502,
Trainer's Signature Title Date 7

| Without preceding “90”

Waters Confidential Template 730001104 VO1

Itis the User's responsibility to verify they are using the current released version
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SKW-MFG4018-00 Rev: A
Temperature and Humidity Monitoring Record
% J Form A
“'Month / Year : Mar 20 “ Equipment Used : Thermohverograph TEQID: 0418
®Floor Level : Sth Floor e ® Serial Number : 5650 " Temperature Range : 18°ct030%¢
“Area_: Production  Calibration due date : 07 Jan 21  Humidity Range : 25%RH to 85%R)
Measurement a 1 Out of Control Record
5 " Within [ o = =3
& [ “ Min. Spec.(Check ok
Date MinTemp. | MaxTemp.| prumidity s | since the st | Temp. | Humidiy | Date | Time =
Cheek [ ime cheek | € co | corm recording) | (O [ %R | Out [ Our @ Name / Sigature NCMR
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