IQ-02 TEST SCRIP DATA SHEET:

AMBIENT TEMPERATURE: 21 ºC to 22 ºC
Verified using

THERMOHYGROGRAPH

Serial Number : 5650

EQID : 0418EQID : 0418

Calibration Due Date: 07-Jan-2021

RELATIVE HUMIDITY: 51 % to 55 %
Verified using

THERMOHYGROGRAPH

Serial Number : 5650

EQID : 0418

Calibration Due Date: 07-Jan-2021

ELECTRICAL SUPPLY :

Verified using :

FLUKE 110 TRUE RMS MULTIMETER

Serial Number: 85730619

EQID: 0245

Calibration Due Date: 16-Dec-2020

Minimum 4 dedicated outlets

ELECTRICAL OUTLET

Computer Work

Station (AC Line CPU)

LCD Monitor (AC Line_

TE82526 Pump

Tray Tester (AC Line)

AND HR 200 Electronics Balance  (AC Line)

NETWORK AVAILABLE: Minimum of 1 dedicated network

Blue Sticker : Dedicated network for Pump Tray Area

WS-9 is a dedicated network

WORK SPACE: The system has sufficient room for the test station, UUT and service

TE82526 Pump

Tray Tester

TEST STATION BENCH :   The test station bench is large and strong enough to support the station

Test Station Bench

IQ-04: Test Software and Instrument Firmware Version

IQ-04: Test Software and Instrument Firmware Version

IQ-05 Test Script Data Sheet:

Station does not have Pinch Points/ sharp edges

Glass Bottles secured inside the container

Glass Bottle is secured by plastic container

Waste Bottles secured inside the container

Wastes Bottle / Glass  secured inside the container