For the 2690 Pump Tray Tester (TE82526)

PREPARED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Luo HuaZhen | Senior Test Engineer |  |  |

The “Prepared By” signature denotes that this document accurately describes the qualification activities.

REVIEWED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Mark Jussuame | Senior Test Engineer |  |  |

The “Reviewed By” signature denotes agreement with the description of the qualification activities.

APPROVED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Kah Chin TAN | Senior Quality Assurance Manager |  |  |

The “Approved By” signature denotes that this document has been reviewed and approved for overall quality and completeness.

REVISION HISTORY

| Version: | 00 | Originator: | Luo HuaZhen | Change ID: | N/A |
| --- | --- | --- | --- | --- | --- |
| Summary of Change: First document for approval |  |  |  |  |  |

# PURPOSE

The purpose of this Installation Qualification Report is to detail the activities associated with the execution of Test Specification and Procedure, Alliance Pumptray WAT210700, Rev 09 for the TE82526, the 2690 Pump Tray Tester.

# SYSTEM BACKGROUND

The TE82526, the 2690 Pump Tray Tester, facilitates the functional test of the following part numbers: 279000419, 279000524, 289000526, WAT270463, WAT271078, and the equivalent spare part numbers. Upon completion of all testing, a record is automatically generated and printed for the operator to collect.

This 2690 Pump Tray Tester TE82526 at the Flextronics Contract Manufacturing Site in Singapore is being retrofitted with new valves in the test box and new computer; This computer will run on Windows 10. The replacement valves are not identical but are functionally equivalent. This is a hardware change that does not change the way the system controls or measures a test specification or PSC result. The custom test software is being installed on the new PC; the change in operating system does not required any changes to the test software. This system will be delivered and installed with the new valve and PC running Windows 10.

# SCOPE

This Installation Qualification Report details the Qualification activities undertaken for the TE82526, the 2690 Pump Tray Tester,  based on the requirements of the executed protocol, the Installation Qualification Protocol, 730004135 Ver. 03 as detailed in TE82526-29JAN2020-VAF.

# SUMMARY OF RESULTS – STATEMENT OF FITNESS FOR INTENDED USE

The TE82526, the 2690 Pump Tray Tester was tested in accordance with the Installation Qualification Protocol, 730004135 Ver. 03.  All Test Scripts were performed, and all objective evidence requirements have been met.

Waters now has the necessary documentation in place to ensure that the TE82526, the 2690 Pump Tray Tester complies with Waters’ requirements, the latest relevant quality system regulations, and cGMP practices.

Following the approval of this Qualification Report, the TE82526, the 2690 Pump Tray Tester can be considered fully qualified and fit for intended use.

# DEFINITIONS / ABBREVIATIONS

The following definitions are used in this document:

| Term | Explanation |
| --- | --- |
| cGMP | Current Good Manufacturing Practice |
| IQ | Installation Qualification |
| IQR | Installation Qualification Report |
| TSP | Test Specification and Procedure |
| CAL/PM | Calibration/Preventive Maintenance |
| S/N | Serial Number |
| P/N | Part Number |
| VAF | Validation Assessment Form |
| Hz | Hertz |
| N/A | Not Applicable |
| SAP | Enterprise Resource Planning Software |
| UUT | Unit Under Test |
| VAC | Volts Alternating Current |
| EQID | Equipment Qualification Identification |
| ANSI | American National Standards Institute |
| PSC | Production Support Checks |

# TRAINING

The operator of the TE82526, the 2690 Pump Tray Tester was trained according to the Test Specification and Procedure, Alliance Pumptray WAT210700, Rev 09. The executor of this qualification was trained to the Installation Qualification Protocol, 730004135 Ver. 03. Both Training records are provided as Attachments 7.1 and 7.2, respectively.

# QUALIFICATION Results

## Test Script IQ-01: Associated Documentation Verification was performed according to Installation Qualification Protocol 730004135, Ver. 03.  The completed Test Script Data Sheet is provided as Attachment 7.3.

#### The Test Specification and Procedure, Alliance Pumptray WAT210700, Rev 09 was confirmed to be in SAP with a status of Release General, as evidenced below:

## Test Script IQ-02: Site Verification was performed according to Installation Qualification Protocol 730004135, Ver. 03.  The completed Test Script Data Sheet is provided as Attachment 7.4.

#### The ambient temperature and humidity were measured using Temperature/Humidity Meter and were determined to be within specification as defined by the Performance Specification, Pump Mechanics Assembly, WAT270698. The calibration record is provided as Attachment 7.5.

#### The electrical supply was measured using a Multimeter and was determined to be within specification as defined by the Performance Specification, Pump Mechanics Assembly, WAT270698.  The calibration record is provided as Attachment 7.6.

#### The Test Station was confirmed to have more than the four (4) dedicated electrical outlets and minimum of 1 dedicated network required. Picture of dedicated electrical outlets and network port is provided as Attachment 7.7.

#### The space and test bench of the TE82526, the 2690 Pump Tray Tester were confirmed to have adequate space for testing and service.  The bench was confirmed to be large and strong enough to support all test equipment.  Picture of the Test Station work area is provided as Attachment 7.8.

## Test Script IQ-03: Instrumentation Calibration Verification was performed according to Installation Qualification Protocol 730004135, Ver. 03.  The completed script is provided as Attachment 7.9.

#### The Electronics Weigh Balance of the TE82526, the 2690 Pump Tray Tester was confirmed to have had calibration within this time frame. The most recent calibration certificate is provided as Attachment 7.10.

#### The TE82526, the 2690 Pump Tray Tester was confirmed to be in the Preventive Maintenance Master Schedule with a six (6) month maintenance cycle.  A screen capture from the database is provided as Attachment 7.11.

## Test Script IQ-04: Software and Instrument Firmware Version was performed according to Installation Qualification Protocol 730004135, Ver. 03.  The completed Test Script Data Sheet is provided as Attachment 7.12.

#### The Test Equipment Software version listed in the TSP is Test Software 1.0 and was verified to match the version listed on the test equipment display.  The objective evidence was provided as Attachment 7.13.

#### This test station does not download firmware.

## Test Script IQ-05: General Safety and Alarms was performed according to Installation Qualification Protocol 730004135, Ver. 03.  The completed Test Script Data Sheet is provided as Attachment 7.14.

#### It was confirmed that there were no mechanical safety pinch points or sharp edges. Picture of tester station safety check is provided as Attachment 7.15

## Test Script IQ-06: Functional Verification was performed according to Installation Qualification Protocol 730004135, Ver. 03.  The completed Test Script Data Sheet is provided as Attachment 7.16.

#### One (1) UUTs was tested using the TE82526, the 2690 Pump Tray Tester with all tests passing.  The test record is provided as Attachment 7.17.

## This table provides a summary of each test script as performed according to 730004135, Ver. 03 along with the attachment location of all objective evidence.  The table also indicates any deviation taken from the scripts during qualification.

| Script or Section | Pass/Fail | Objective Evidence | Deviations |
| --- | --- | --- | --- |
| IQ-01 | Pass | Attachment 7.3 | None |
| IQ-02 | Pass | Attachments 7.4, 7.5, 7.6, 7.7, and 7.8 | None |
| IQ-03 | Pass | Attachments 7.9, 7.10, and 7.11 | None |
| IQ-04 | Pass | Attachments 7.12 and 7.13 | None |
| IQ-05 | Pass | Attachments 7.14 and 7.15 | None |
| IQ-06 | Pass | Attachment 7.16, and 7.17 | None |

# ATTACHMENTS

## Completed Operator Training Record WAT210700, Rev 09

## Completed Executor Training Record 730004135, Ver. 03

## Completed Test Script IQ-01: Associated Document Verification

## Completed Test Script IQ-02: Site Verification

## Objective Evidence IQ-02: Temperature/Humidity Meter Calibration Record and Measurement Record

## Objective Evidence IQ-02: Multimeter Calibration Record and Measurement Record

## Objective Evidence IQ-02: Picture of dedicated electrical outlets and network port

## Objective Evidence IQ-02: Picture of Test Station Work Area

## Completed Test Script IQ-03: Instrument Calibration Verification

## Objective Evidence IQ-03: Electronics Weigh Balance Calibration Record

## Objective Evidence IQ-03: Screen Capture of Preventive Maintenance Master Schedule

## Completed Test Script IQ-04: Test Software and Instrument Firmware Version

## Objective Evidence IQ-04: Test Software version on Cover Page of TSP WAT210700 and on Test Equipment

## Completed Test Script IQ-05: General Safety and Alarms

## Objective Evidence IQ-05: Picture of tester station safety check

## Completed Test Script IQ-06: Functional Verification

## Objective Evidence IQ-06: Test Records for UUT

# DEVIATION FROM THE QUALIFICATION PROTOCOL

There were no deviations generated during the execution of the Installation Qualification Protocol 730004135, Ver. 03.

# References

The current approved versions of the following documents are referenced within this report:

| Document Number | Document title |
| --- | --- |
| 730004135, Ver. 03 | Installation Qualification Protocol |
| WAT210700, Rev 09 | TEST SPECIFICATION AND PROCEDURE, ALLIANCE PUMPTRAY |
| WAT270698, Rev A | PERFORMANCE SPECIFICATION, Pump Mechanics Assembly |
| WAT046538, Ver. 12 | Generic Electrical Safety Test Procedure |
| ANSI C84.1 | Voltage Tolerance Standard |
| TE82526-29JAN2020-VAF | Validation Assessment Form for the 2690 Pump Tray Tester |