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IQ-02 TEST SCRIP DATA SHEET: 

AMBIENT TEMPERATURE: 21 ºC to 22 ºC
Verified using 
THERMOHYGROGRAPH
Serial Number : 5650
EQID : 0418EQID : 0418
Calibration Due Date: 07-Jan-2021
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RELATIVE HUMIDITY: 56 % to 58 %
Verified using 
THERMOHYGROGRAPH
Serial Number : 5650
EQID : 0418EQID : 0418
Calibration Due Date: 07-Jan-2021
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ELECTRICAL SUPPLY : Minimum 4 dedicated outlets
Verified using : 
FLUKE 110 TRUE RMS MULTIMETER
Serial Number: 85730619
EQID: 0245
Calibration Due Date: 16-Dec-2020
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ELECTRICAL OUTLET 
TE82526-2 Pump 
Tray Tester (AC Line)

Computer  Work
Station (AC Line CPU)
LCD Monitor (AC Line_




[image: ]AND HR 200 Electronics Balance  (AC Line)
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NETWORK AVAILABLE : Minimum of 1 dedicated network
Blue Sticker : WS-8 is a dedicated network
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WS-8 is a dedicated network






















WORK SPACE: The system has sufficient room for the test station, UUT and service
TE82526-2 Pump 
Tray Tester
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  TEST STATION BENCH :   The test station bench is large and strong enough to support the station                                                      [image: ]Test Station Bench




IQ-04 Software and Instrument Firmware Version 

[bookmark: _GoBack][image: ]


IQ-05 Test Script Data Sheet:

Station does not have Pinch Points/ sharp edges
Glass Bottle is secured by plastic container
Glass Bottles secured inside the container

                [image: ]Secure Glass Bottle



Wastes Bottle / Glass  secured inside the container







IQ-06 Test Script Data Sheet:
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Waters Corporation
SUBASSEMBLY TEST RECORD

Part Number 289000526 Description: _ Allance Pump Tray
Tested By 26154 Date: 121212020 3:12:83 pm
Serial Number 06LOHO71 Control Number: NA
Software Revision: NA Factory Order NA |
Seal Wash Functonal Test PASSI FAILL | see note # PASS

Accumulator Decay Test

Total decay (<150 psi) 42 Sl PASSX FAILL| seenote# PASS
Compress Volume (<20 %) - Actual: 17 % PASSX| FAIL[| seenote# PASS
Primary Decay Test
Total decay (<150 psi) - Actual: 30 PSI  PASSX| FAL | scenote# PASS

Compress Volume (<20 %) - Actual: 10 %  PASSX| FAILL| seenote# paSS
Flow Rate Test [Heasursa Fow: 0,098 =]

(<= +-1%)-Actal; 022 %ERROR  PASSX| FAIL[ | seencted PASS
Ripple Test -

(¢=25%) - Actual: 1.4 % PSS FAIL | seenole# __ PASS

Pump Tray Result: PASS X< FAIL Signed: 246154 JOEL Date: (™ fies | 202
Notes: 1

216000425 €
Pump POB SN 11012046 1 Aecum Xducar SIN;_ 295 842 Prim. Kaucer SN:_295 407

i Test Rcord sy 428 cortined i i ar propreany o Woters Corparaton 37 must b (et by Waters for the ime |
et e n s Qualty Records Retention Py, I ryof theinidal ests nees above ;2 ot st be refeenced tht
el the werk parformed  orrect st seciic defceny,An cveral fallrendiates  specl crcumstance n wich cse
 complte ettt b done. A second o 5 4quie 10 decument ta esing th it fom st b atached and recaine,





image15.png
Waters WORK INSTRUGTION

THE SCIENCE OF WHAT'S POSSIBLE® Installation Qualification Protocol
Document # 730004135 Version #03 Page 708

Appendix A: TRAINING RECORD

TEST SPECIFICATION AND PROCEDURE,
WAT210700 | Document itie PTGl

:

., Ny 10 re 20x)

Printed Name Signature Date

and Understosd (1 seltirained — If no self-rainec his sacton can be deletes

 signing you acknowedge that you have reviewed, understand, and agree fo comply with the polcies and practices of the
rocedure defined above.
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SN
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Asiachment Pege __of

Waters Corfderiial Termplate 730000483 V08
1 heresponaiity of ihe sor o erty thy ar using the currnt relased vasion.
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B Revl | Type Document Vs Part Description Status text Current Ve Vaiid From |Vaid To | Change No. |Long name
05 MSC WAT210700 05 000 TEST PROC., 2690 PUMP TRAY TESTER RELEASE GENERAL 03.07.2002  31.12.9999 9000280  PLM MIGRA,
00 TSP WAT210700 00 000 TEST SPEC/PROC,M2690 PUMP TRAY WBW/O, RELEASE GENERAL 11.08.1998  28.10.1998 ECO3831  Production
01 TSP WAT210700 01 000 TEST SPEC/PROC,M2690 PUMP TRAY W8W/O, RELEASE GENERAL 29101998  23.02.1999 EC04099  Production
02 TSP WAT210700 02 000 TEST SPEC/PROC,M2690 PUMP TRAY WBWW/O, RELEASE GENERAL 24.02.1999  14.06.2000 ECO4523  Production
03 TSP WAT210700 03 000 TEST SPEC/PROC,M2690 PUMP TRAY WBW/O, RELEASE GENERAL 15.06.2000  21.02.2001 ECO06593 Production
04 TSP WAT210700 04 000 TEST SPEC/PROC,M2690 PUMP TRAY WBWW/O, RELEASE GENERAL 22.02.2001  20.06.2001 EC007629 Rosely Dow,
05 TSP WAT210700 05 000 TEST SPEC/PROC,M2690 PUMP TRAY W8W/O, RELEASE GENERAL 21.06.2001  02.06.2003 ECO08175 Rosely Dow,
06 TSP WAT210700 06 000 TEST SPEC/PROC,M2690 PUMP TRAY WBW/O, RELEASE GENERAL 03.06.2003  30.11.2014 9104161  Wendy Har,
07 TSP WAT210700 07 000 TEST SPEC/PROC,M2690 PUMP TRAY WBW/O, RELEASE GENERAL 01122014 08.06.2015 9193357  Michael Cot,

WAT210700 000 TEST SPEC/PROC,M2690 PUMP TRAY WEV/O_ RELEASE GENERAL 09.06.2015  08.05.2016 9198230 Michael Cot.
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