# INSTALLATION QUALIFICATION REPORT

For the 2489 Acquity TUV Pre-test Download Station (TE82964-1)

PREPARED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Emma Pierce | Test Validation Engineer |  |  |

REVIEWED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Mark Jussaume | Senior Test Engineer |  |  |

APPROVED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Andre Kopoyan | Senior Operations Quality Engineer |  |  |

The “Prepared By” signature denotes that this document accurately describes the qualification activities.

The “Reviewed By” signature[s] denotes agreement with the description of the qualification activities.

The “Approved By” signature denotes that this document has been reviewed and approved for overall quality and completeness.

REVISION HISTORY

| Version | Author | Summary of change |
| --- | --- | --- |
| 00 | Emma Pierce | First document for approval |

# Purpose

##### The 2489 Acquity TUV Pre-test Download Station (TE82964-1) is located in the Flextronix Contractor Manufacturer Site in Singapore.  The system is designed to download instrument firmware and assign the instrument serial number for the following part numbers: 186017002, 186248900, and 186002487.  The UUT is then tested on a TUV Final Test Station (TE82903).  The final test stations at the Flextronix Site used for these part numbers are TE82903-6, TE82903-8, TE82903-9, and TE82903-10.

##### This Installation Qualification is being performed as required by Project CYAN.  This change impacts only the firmware file being loaded to the UUT.  The new firmware is verified to download to the UUT correctly via Final Test Records from the TUV Final Test Stations (TE82903-6, TE82903-8, TE82903-9, and TE82903-10).

# Scope of Report

This Installation Qualification Report details the requalification activities undertaken for the 2489 Acquity TUV Pre-test Download Station (TE82964-1) based on the requirements of the Installation Qualification Procedure, 730004135 Ver. 00 and TE82903-13JUL2017-VAF.

This Installation Qualification Report defines the documents used to qualify the 2489 Acquity TUV Pre-test Download Station (TE82964-1).

# Summary of Results - Statement of Fitness for Intended Use

The 2489 Acquity TUV Pre-test Download Station was tested in accordance with the Installation Qualification Procedure, 730004135 Ver. 00.  All Test Scripts were performed and all objective evidence requirements have been met.  This download station has been classified as Fit for Use before the execution of this report.

Waters now has the necessary documentation in place to ensure that the 2489 Acquity TUV Pre-test Download Station complies with Waters’ requirements, the latest relevant quality system regulations, and cGMP practices.

Following the approval of this Qualification Report, the 2489 Acquity TUV Pre-test Download Station (TE82964-1) can be considered fully qualified and fit for intended use.

# Glossary

The following definitions are used in this document:

| Term | Explanation |
| --- | --- |
| cGMP | Current Good Manufacturing Practice |
| IQ | Installation Qualification |
| IQR | Installation Qualification Report |
| TSP | Test Specification and Procedure |
| CAL/PM | Calibration/Preventive Maintenance |
| UUT | Unit Under Test |
| P/N | Part Number |
| N/A | Not Applicable |
| ECR | Engineering Change Request |
| VAF | Validation Assessment Form |
| TEQ | Test Equipment Qualification |
| TEQR | Test Equipment Qualification Report |
| TUV | Tunable UV/Vis Absorbance Detector |

# Training

##### The operator of the TUV Final Test Station was trained according to the Acquity Arc TUV Final Test Procedure, 735003545, Ver. 01, the TUV Detectors Final Test Procedure, 735001356, Ver. 11, and the 2489 UV/Visible Detector Test Procedure, 735003877, Ver. 01.  These training records are provided as Attachment 7.1.  The executor of this requalification was trained according to the Installation Qualification Procedure, 7350004135, Ver. 00. This training record is provided as Attachment 7.2.

# Qualification Results

## Test Script IQ-01: Associated Documentation Verification was not required to be performed as part of this test equipment qualification according to TE82903-13JUL2017-VAF.

#### This system is classified as fit for use.

## Test Script IQ-02: Site Verification was not required to be performed as part of this test equipment qualification according to TE82903-13JUL2017-VAF.

#### This system is classified as fit for use.

## Test Script IQ-03: Instrumentation Calibration Verification was not required to be performed according to TE82903-13JUL2017-VAF.

#### This system is classified as fit for use.

## Test Script IQ-04: Software and Instrument Firmware Version was performed according to Installation Qualification Procedure 730004135, Ver. 00.  The completed Test Script Data Sheet is provided as Attachment 7.3.

#### The 2489 Acquity TUV Pre-test Download Station (TE82964-1) does not have any custom software to be verified.

#### The firmware version listed in ECR #9215283 is version 1.69.2682 all part numbers.  The correct firmware download for each part number were verified to match the versions listed on the Final Test Records from each TUV Final Test Station.  The Final Test Records for TE82903-6 are provided as Attachment 7.4, the Final Test Records for TE82903-8 are provided as Attachment 7.5, the Final Test Records for TE82903-9 are provided as Attachment 7.6, and the Final Test Records for TE82903-10 are provided as Attachment 7.7.

#### P/N: 289024872 is a sub-assembly of P/N: 186002487 and is tested as P/N: 186002487.  The upper level part (P/N: 186002487) is the packaging assembly that contains P/N: 289024872 along with the packing materials, labels, etc. required for shipping.  No changes are made to P/N: 289024782 after it is tested.

## Test Script IQ-05: General Safety and Alarms was not required to be performed as part of this test equipment qualification according to TE82903-13JUL2017-VAF.

#### This system is classified as fit for use.

## Test Script IQ-06: Functional Verification was not required to be performed as part of this test equipment qualification according to TE82903-13JUL2017-VAF.

#### This system is classified as fit for use.

## This table provides a summary of each test script as performed according to 7300004135 Ver. 00 along with the attachment location of all objective evidence.  The table also indicates any deviation taken from the scripts during qualification.

| Script or Section | Pass/Fail | Objective Evidence | Deviations |
| --- | --- | --- | --- |
| IQ-04 | Pass | Attachments 7.3, 7.4, 7.5, 7.6, and 7.7 | None |

# Attachments

## Completed Operator Training Record 735003545, Ver. 01

## Completed Executor Training Record 730004135, Ver. 00

## Completed Test Script IQ-04: Software and Instrument Firmware

## Objective Evidence IQ-04: Instrument Final Test Records for TE82903-6

## Objective Evidence IQ-04: Instrument Final Test Records for TE82903-8

## Objective Evidence IQ-04: Instrument Final Test Records for TE82903-9

## Objective Evidence IQ-04: Instrument Final Test Records for TE82903-10

# Deviation from Test Equipment Qualification

There were no deviations generated during the execution of the Installation Qualification Procedure 730004135 Ver. 00.

# References

The current approved versions of the following documents are referenced within this report:

| Document Number | Document title |
| --- | --- |
| 730004135, Ver. 00 | ###### Installation Qualification Procedure |
| 735003545, Ver. 01 | ###### Acquity Arc TUV Final Test Procedure |
| 735001356, Ver. 11 | ###### TUV Detectors Final Test Procedure |
| 735003877, Ver. 01 | ###### 2489 UV/Visible Detector Test Procedure |
| ECR #9215283 | ###### Engineering Change Request for Project CYAN |
| ###### TE82903-13JUL2017-VAF | ###### Validation Assessment Form for Project CYAN Implementation |