# installation qualification REPORT

For the TUV DOWNLOAD / PRETEST STATION (TE82964-1)

PREPARED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Luo HuaZhen | Senior Test Engineer |  |  |

The “Prepared By” signature denotes that this document accurately describes the qualification activities.

REVIEWED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Eric Turner | Principal Test Engineer |  |  |

The “Reviewed By” signature denotes agreement with the description of the qualification activities.

APPROVED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Kah Chin TAN | Senior Quality Assurance Manager |  |  |

The “Approved By” signature denotes that this document has been reviewed and approved for overall quality and completeness.

REVISION HISTORY

| Version: | 02 | Originator: | Luo HuaZhen | Change ID: | ECR9236194 |
| --- | --- | --- | --- | --- | --- |
| Summary of Change: Add new ACQUITY PREMIER TUV DETECTOR part number 186018006, 800000534 |  |  |  |  |  |

# PURPOSE

The purpose of this Installation Qualification Report is to detail the activities associated with the execution of Test Specification and Procedure, ACQUITY PREMIER TUV DETECTOR 735005108, Rev A for the TE82964-1, the TUV DOWNLOAD / PRETEST STATION.

## SYSTEM BACKGROUND

The TE82964-1, the TUV DOWNLOAD / PRETEST STATION, is designed to download instrument firmware and assign the instrument serial number for 4289, ACQUITY TUV Detectors and ACQUITY PREMEIER TUV Detectors. As per Test Specification and Procedure, ACQUITY PREMIER TUV DETECTOR 735005108, the station can support for the following part numbers: 186018006 and 800000534.

The TE82964-1, the TUV DOWNLOAD / PRETEST STATION is located at the Flex Contract Manufacturing Site in Singapore. The UUT is connected to a custom test box that simulates a leak and uses HyperTerminal to download a Serial Number to the UUT. The system does not have custom software, nor does it print a Test Record.

# SCOPE

This Installation Qualification Report details the Qualification activities undertaken for the TE82964-1, the TUV DOWNLOAD / PRETEST STATION,  based on the requirements of the executed protocol, the Installation Qualification Protocol, 730004135 Ver. 03 as detailed in TE82964-05AUG20-VAF.

# SUMMARY OF RESULTS – STATEMENT OF FITNESS FOR INTENDED USE

The TE82964-1, the TUV DOWNLOAD / PRETEST STATION was tested in accordance with the Installation Qualification Protocol, 730004135 Ver. 03.  All Test Scripts were performed, and all objective evidence requirements have been met.

Waters now has the necessary documentation in place to ensure that the TE82964-1, the TUV DOWNLOAD / PRETEST STATION complies with Waters’ requirements, the latest relevant quality system regulations, and cGMP practices.

Following the approval of this Qualification Report, the TE82964-1, the TUV DOWNLOAD / PRETEST STATION can be considered fully qualified and fit for intended use.

# DEFINITIONS / ABBREVIATIONS

The following definitions are used in this document:

| Term | Explanation |
| --- | --- |
| cGMP | Current Good Manufacturing Practice |
| IQ | Installation Qualification |
| IQR | Installation Qualification Report |
| TSP | Test Specification and Procedure |
| S/N | Serial Number |
| P/N | Part Number |
| VAF | Validation Assessment Form |
| N/A | Not Applicable |
| SAP | Enterprise Resource Planning Software |
| UUT | Unit Under Test |
| PSC | Production Support Checks |
| TUV | Tunable UV/Vis Absorbance Detector |
| ECR | Engineering Change Request |
| UPLC | Ultra Performance Liquid Chromatography |

# TRAINING

The operator of the TE82964-1, the TUV DOWNLOAD / PRETEST STATION was trained according to the Test Specification and Procedure, ACQUITY PREMIER TUV DETECTOR 735005108, Rev A. The executor of this qualification was trained to the Installation Qualification Protocol, 730004135 Ver. 03. Both Training records are provided as Attachments 7.1 and 7.2, respectively.

# QUALIFICATION Results

## Test Script IQ-01: Associated Documentation Verification was performed according to Installation Qualification Protocol 730004135, Ver. 03.  The completed Test Script Data Sheet is provided as Attachment 7.3.

#### The Test Specification and Procedure, ACQUITY PREMIER TUV DETECTOR 735005108, Rev A was confirmed to be in SAP with a status of Ready. A screen capture from SAP is provided as Attachment 7.4.

## Test Script IQ-02: Site Verification was not required to be performed as part of this test equipment qualification according to TE82964-05AUG20-VAF.

## Test Script IQ-03: Site Verification was not required to be performed as part of this test equipment qualification according to TE82964-05AUG20-VAF.

## Test Script IQ-04: Site Verification was not required to be performed as part of this test equipment qualification according to TE82964-05AUG20-VAF.

## Test Script IQ-05: Site Verification was not required to be performed as part of this test equipment qualification according to TE82964-05AUG20-VAF.

## Test Script IQ-06: Functional Verification was performed according to Installation Qualification Protocol 730004135, Ver. 03.  The completed Test Script Data Sheet is provided as Attachment 7.5.

#### One (1) UUTs was tested using the TE82964-1, the TUV DOWNLOAD / PRETEST STATION with all tests passing.  The screenshot of passing result is provided as Attachment 7.6.

## This table provides a summary of each test script as performed according to 730004135, Ver. 03 along with the attachment location of all objective evidence.  The table also indicates any deviation taken from the scripts during qualification.

| Script or Section | Pass/Fail | Objective Evidence | Deviations |
| --- | --- | --- | --- |
| IQ-01 | Pass | Attachment 7.3 and 7.4 | None |
| IQ-06 | Pass | Attachments 7.5 and 7.6 | None |

# ATTACHMENTS

## Completed Operator Training Record 735005108, Rev A

## Completed Executor Training Record 730004135, Ver. 03

## Completed Test Script IQ-01: Associated Document Verification

## Objective Evidence IQ-01: Screenshot of SAP

## Completed Test Script IQ-06: Functional Verification

## Objective Evidence IQ-06: Screenshot of passing result on UUT

# DEVIATION FROM THE QUALIFICATION PROTOCOL

There were no deviations generated during the execution of the Installation Qualification Protocol 730004135, Ver. 03.

# References

The current approved versions of the following documents are referenced within this report:

| Document Number | Document title |
| --- | --- |
| 730004135, Ver. 03 | Installation Qualification Protocol |
| 735005108, Rev A | Test Spec and Procedure, ACQUITY PREMIER TUV DETECTOR |
| ECR9236194 | Engineering Change Request for Project ACQUITY PREMIER TUV DETECTOR |
| TE82964-05AUG20-VAF | Validation Assessment Form for the TUV DOWNLOAD / PRETEST STATION |