For the  PCM II Tester (TE82817-2)

PREPARED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Luo HuaZhen | Senior Test Engineer |  |  |

The “Prepared By” signature denotes that this document accurately describes the planned Process Validation activities.

REVIEWED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Matt Brule | Principal Test Engineer |  |  |

The “Reviewed By” signature[s] denotes agreement with the details of the planned testing and/or deliverables to be performed.

APPROVED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Kah Chin Tan | Senior Quality Assurance Manager |  |  |

The “Approved By” signature denotes that this system change risk assessment and resulting planned process validation activities has been approved.

# Purpose

The purpose of the Validation Assessment Form (VAF) is to document the assessment of an engineering change to the Pump Control Module,  PCM II Tester (TE82817-2), and to define the resulting deliverable(s) required to maintain the system process validation.

# Background

The PCM II Tester (TE82817-2), facilitates the functional test of the following part numbers: 186002449, 800000178, 279001290, 800000224. The system downloads firmware and verifies the electrical and mechanical functionality of the unit under test. Upon completion, a Final Test Record is automatically generated and printed for the operator to collect.

# Description of Change

## The computer in use at the PCM II Tester (TE82817-2) located in CEI Singapore has been replaced with a new PC. The new computer uses the same operating system, Windows XP, and there was no change to the custom test software. This change does not affect TE82817-1 located in Milford NPI.

# Risk assessment

## High Risk Assessment

## Please check the appropriate boxes below to determine the high risk assessed qualification or requalification requirements:

| ### Specification | ## No | ## Yes |
| --- | --- | --- |
| Is the change a test specification value being modified and/or added? |  |  |

| ### Software/Firmware | ## No | ## Yes |
| --- | --- | --- |
| Is the change a software modification that affects the way the system controls, and/or measures a test specification and/or PSC result? |  |  |
| Is the change due to firmware downloaded to the UUT by a test system? |  |  |
| Is the change a flash file update? |  |  |
| Is it a change to an existing, or the addition of a new, PCB Test Sequence? |  |  |

| ### Hardware (Mechanical, Electrical, Pneumatic, Hydraulic, etc.) | ## No | ## Yes |
| --- | --- | --- |
| Is the change a mechanical, electrical, pneumatic, and/or hydraulic modification that alters the way the system controls, and/or measures a test specification and/or PSC result? |  |  |

| ### Equipment/Process Moves | ## No | ## Yes |
| --- | --- | --- |
| Is the change a move of the equipment, test method or process, to a different facility? |  |  |

## Medium Risk Assessment

## Please check the appropriate boxes below to determine the medium risk assessed qualification or requalification requirements:

| ### Replacement of Non-Identical | ## No | ## Yes |
| --- | --- | --- |
| Is the change due to a replacement of a system component with a non-identical equivalent? |  |  |

| ### Computer Related or Replacement of Custom PCBs | ## No | ## Yes |
| --- | --- | --- |
| Is the change an update to the PC operating system? |  |  |
| Is the change a computer or printed circuit board replacement? |  |  |

| ### Software/Firmware | ## No | ## Yes |
| --- | --- | --- |
| Is the change a minor software bug fix or a minor improvement (i.e. a part number add)? |  |  |
| Is the change a test system instrument firmware update (i.e. update to a stack component)? |  |  |

| ### Equipment/Process Moves | ## No | ## Yes |
| --- | --- | --- |
| Is the change a move of the equipment, test method or process, within the local site? |  |  |

## This change results in a Medium Risk Assessment and the associated process validation requirements, as detailed in Section 6.0.

# Hazard Assessment:

There are no hazards potentially introduced to the environment, operator, assembler, service technician or end user due to

this change.

# verification, Qualification and/or validation:

| ## Equipment Design Verification (EDV) | ## No | ## Yes |
| --- | --- | --- |
| Based on the Risk Assessment, does an EDV Protocol need to be created or revised due to this change? |  |  |
| Based on the Risk Assessment, does an EDV Protocol need to be executed due to this change? |  |  |
| Comments/Justification An EDV is not required because this change does not affect the way the system controls or measures a specification or PSC result. |  |  |

| ## Installation Qualification (IQ) | ## No | ## Yes |
| --- | --- | --- |
| Based on the Risk Assessment, does an IQ Protocol need to be executed due to this change? |  |  |
| Comments/Justification A functional test will be performed on the PCM II Tester (TE82817-2) located in CEI Singapore to verify the custom test software and instrument firmware versions, as well as to verify the functionality of the system. In lieu of a formal IQ Report, the objective evidence will be attached to the provided Objective Evidence Completion form and submitted for approval. |  |  |

| ## Test Method Validation (TMV) | ## No | ## Yes |
| --- | --- | --- |
| Based on the Risk Assessment, does a TMV need to be executed due to this change? |  |  |
| Comments/Justification TMV is not required because this system does not utilize external measurement equipment and this change does not involve Operator influence. |  |  |

| ## Operational Qualification (OQ) | ## No | ## Yes |
| --- | --- | --- |
| Based on the Risk Assessment, does an OQ Protocol need to be created or revised due to this change? |  |  |
| Based on the Risk Assessment, does an OQ need to be executed due to this change? (Justify partial execution in the Comments/Justification section below.) |  |  |
| Comments/Justification An OQ is not required because 100% of critical to quality attributes are verified. |  |  |

| ## Performance Qualification (PQ) | ## No | ## Yes |
| --- | --- | --- |
| Based on the Risk Assessment, does a PQ Protocol need to be created or revised due to this change? |  |  |
| Based on the Risk Assessment, does a PQ need to be executed due to this change? (Justify partial execution in the Comments/Justification section below.) |  |  |
| Comments/Justification A PQ is not required because 100% of critical to quality attributes are verified. |  |  |

## Electromagnetic Interference (EMI) Assessment:

This change only affects test equipment; therefore, there is no impact to product EMI.

# IMPLEMENTATION ASSESSMENT:

- There is no impact to documentation, calibration, or facilities/utilities.

- There is no impact to any other qualified, verified, or validated process, test method, or equipment.

- The production schedule for this system is currently stopped awaiting validation.

- This change will be implemented by Test Engineering; the qualification activities will be executed by a Test Engineer.

- No additional training be required.

- Product used for qualification or tested prior to approval of qualification activities shall be quarantined.

- Should any predefined deliverable to unable to be obtained the deviation should be addressed as an additional attachment to the Objective Evidence Completion Form.

# Validation Assessment
Objective Evidence Completion

For the  PCM II Tester (TE82817-2)

PREPARED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Luo HuaZhen | Senior Test Engineer |  |  |

The “Prepared By” signature denotes that the document accurately provides the required objective evidence.

REVIEWED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Matt Brule | Principal Test Engineer |  |  |

The “Reviewed By” signature[s] denotes agreement that the required objective evidence is provided and accurate.

APPROVED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Kah Chin Tan | Senior Quality Assurance Manager |  |  |

The “Approved By” signature denotes that the objective evidence has been reviewed and approved for overall quality and completeness.

ATTACHMENTS:

Attachment 1: Cover sheet of the test procedure listing the custom test software version

Attachment 2: BOM for the unit under test listing the required firmware version

Attachment 3: Operator training record

Attachment 4: Passing Test Record